## Instructions for Use
AAFS Standards Resources and Training
Instructions for Use of this Checklist Basic information on use of the checklist is included in select Column Headers as a Comment. Beyond this, use the link above to the AAFS Standards Resources and Training web page to access written instructions, an on-demand training webinar and instructional videos.
NOTE: A Forensic Science Service Provider can use space on this tab to provide agency specific instructions to their staff that will complete this checklist.

## ANSI ASTM E2926-17
ANSI/ASTM E2926-17
Standard Test Method for Forensic Comparison of Glass Using Micro X-ray Fluorescence (µ-XRF) Spectrometry
Discipline: | Trace Evidence
Checklist version: | V1
Standard Section | Section or Clause Number | Clause Type | Clause Wording | FSSP Objective Evidence Document(s) or Records(s) | Implementation | Implementation Status | Reason for Less than Full Implementation | Implementation Plan/Other Notes | Date Implemented or Implementation Timeline Date | Auditing | Auditor Objective Evidence | Auditor Notes | Audit Status | Audit - Opportunity for Improvement | Audit - Nonconformance | Resolution of Nonconformance
Interferences | 5 | Section Title
Interferences | 5.1 | Requirement | Peak overlaps occur in various regions of the EDS spectrum. In glass, such interferences include the overlap of characteristic X-ray lines (for example, Ti K-series and Ba L-series), sum peaks, primary X-ray source excitation peaks (for example, Rh), and escape peaks. In general, automated deconvolution algorithms are included in data processing software that adequately address such overlaps. EDS spectra shall be manually inspected to ensure that potential peak overlaps are considered and addressed.
Apparatus | 6 | Section Title
Apparatus | 6.1 | Requirement | A µ-XRF spectrometer with an EDS detector is employed. Most commercial-grade µ-XRF systems with EDS detectors should be adequate for forensic analysis of glass. The µ-XRF system must, however, meet the following performance specifications:
Apparatus | 6.1.1 | Requirement | The spot size(s) must be within the range(s) of approximately 10 µm to 2 mm; the spot size used may be adjustable to different sizing for instruments with appropriate optics.
Apparatus | 6.1.2 | Requirement | The instrument must be capable of operating at an accelerating voltage of 35 kV or greater.
Apparatus | 6.1.3 | Requirement | The EDS detector must be capable of a resolution that is typically less than 180 eV, measured as the full width at half the maximum height of the Mn Kα peak; better resolutions will provide improved discrimination of adjacent or overlapping peaks, or both.
Apparatus | 6.1.4 | Requirement | A calibrated, scaled display of energy units (keV) and the ability to identify and label X-ray lines is required for the EDS system.
Apparatus | 6.2 | Requirement | Energy Calibration Material Capable of calibrating the EDS detector at both the low (<2 keV) and high (>6 keV) X-ray spectral regions.
Apparatus | 6.3 | Requirement | An X-ray source that does not yield significant spectral interferences with the characteristic X-ray lines for the elements typically found in glass is required. Several X-ray sources are available; a rhodium X-ray source is preferred for appropriate excitation energy and minimal spectral interferences for elements in glass. Other X-ray sources such as Mo X-ray tubes cause interferences with discriminating elements such as Zr.
Apparatus | 6.4 | Requirement | A vacuum sample chamber, sample stage, and visualization system are required.
Apparatus | 6.5 | Requirement | The sample holder, sample support film, and mounting material (for example, adhesive with low trace elements) must prevent background interferences.
Hazards | 7 | Section Title
Hazards | 7.1 | Requirement | The X-ray sources emit radiation when energized. For operator safety, appropriate shielding and safety interlocks must be in place and operational.
Calibration And Standardization | 8 | Section Title
Apparatus | 8.1 | Requirement | The instrument must be optimized as in accordance with manufacturer’s instruction.
Energy Calibration | 8.1.1 | Requirement | Calibrate the X-ray energy scale to characteristic X-ray emission lines by either measuring the centroid energy of a low- (<2 keV) and high- (>6 keV) energy peak or by using software provided by the instrument manufacturer. For example, the aluminum (1.486 keV) and copper (Cu) (8.040 keV) Kα-X-ray energy lines may be used.
Stage Calibration | 8.1.2 | Requirement | For automated or multiple point analysis, initialize the stage position to assure that the stage coordinates accurately reflect the stage position.
Optical Alignment | 8.1.3 | Section Title
Optical Alignment | 8.1.3.1 | Requirement | Align X-ray optics to obtain the maximum count rate.
Optical Alignment | 8.1.3.2 | Requirement | Align visualization optics to ensure that the visual target area coincides with the X-ray beam position.
Spot Size Measurement | 8.1.4 | Requirement | Determine spot size of the X-ray beam at the focal point of the visualization optics. For instruments with continuous variable spot size options, determine the spot size at multiple settings and interpolate the others.
Reference Materials | 8.1.5 | Requirement | Analyze a glass certified reference material (CRM) (for example, NIST SRM 1831) to verify the calibration of X-ray energy lines for elements present in glass and determine if the instrument response is within acceptable limits. Measure this glass CRM using the same analysis parameters as the glass specimens. Use this reference glass sample to normalize element ratios for interlaboratory comparisons, intralaboratory data collection from different analytical runs, and databasing applications to improve precision.
Blanks | 8.1.6 | Requirement | Collect a spectrum of a specimen devoid of elements having an atomic number of 11 or greater, such as the plastic stage plate or an area of the support material having no glass present. Record any system peaks present for future reference.
Quality Assurance | 8.2 | Section Title
Quality Assurance | 8.2.1 | Requirement | The performance of the instrument must be monitored routinely and the frequency and tolerances should be set by each laboratory.
Quality Assurance | 8.2.1.1 | Requirement | Check the system calibration prior to the performance of an analysis.
Quality Assurance | 8.2.1.2 | Requirement | Check the performance of the X-ray source using a known element standard (for example, Cu). Maximum counts for the system should be obtained utilizing system operating parameters established by the laboratory. Maximum counts should not show appreciable drift from acceptable parameters established by the laboratory or analyst for this procedure (10 % tolerance is recommended).
Quality Assurance | 8.2.2 | Requirement | Demonstrate that Ti and Sr have LOD in a soda-lime glass matrix of 75 ppm or less (as described in 11.1) for the instrumental parameters used for collection of spectra from the glass specimens. NIST SRM 1831 is a suitable sample for this purpose.
Procedure | 9 | Section Title
Specimen Preparation | 9.1 | Section Title
Specimen Preparation | 9.1.1 | Requirement | Examine glass fragments using stereomicroscopy to determine an appropriate preparation method for the specimen.
Specimen Preparation | 9.1.2 | Requirement | If necessary, clean the specimen to remove any surface contamination. Cleaning may include washing specimens with soap and water, with or without ultrasonication, and rinsing in deionized water, followed by rinsing in acetone, methanol, or ethanol, and drying. Soaking in various concentrations of nitric acid for 30 minutes or longer, rinsing with deionized water and ethanol, and drying prior to analysis removes most surface contamination without affecting the measured concentrations of elements inherent in the glass. However, the use of nitric acid may remove any surface coating that may be present.
Specimen Preparation | 9.1.3 | Requirement | Mount the specimen for analysis.
Specimen Preparation | 9.1.3.1 | Informational Text | The specimen mounting technique depends on the sample size and shape, beam size, X-ray fluorescence spectrometer chamber design and purpose of the examination.
Specimen Preparation | 9.1.3.2 | Requirement | Raise specimens off the surface of the stage for analysis using an X-ray transparent sample holder or supportive X-ray film, or both. This positioning reduces X-ray scatter off of the surface of the stage and, hence, improves sample signal-to-noise. Because analysis is performed under vacuum, ensure that specimens retain their position on the sample holder by securing with adhesive. Prior to analysis, analyze a small amount of the adhesive to determine the presence of any elements that could interfere with those in the specimen. When small amounts of adhesive are used and beam overspill (X-ray beam extending beyond the perimeter of the specimen) is avoided, little to no interference from the adhesive will be observed.
Specimen Preparation | 9.1.3.3 | Requirement | Position specimens to present as flat a surface as possible to the impinging excitation X-ray beam. If necessary, use a small amount of adhesive to facilitate this positioning.
Specimen Preparation | 9.1.3.4 | Recommendation | For comparisons, glass specimen should be of similar size, shape, and thickness to each other. For full thickness fragments of float glass, comparisons should be made between similar surface types (for example, non-float surface to non-float surface).
Specimen Preparation | 9.1.4 | Requirement | Place sample(s) in the instrument’s analysis chamber. For automated multiple point analyses, it may be necessary to secure the sample/sample holder to the instrument stage.
Specimen Preparation | 9.1.5 | Requirement | Evacuate the chamber; samples should be run under vacuum.
Specimen Preparation | 9.1.6 | Requirement | Target specimen areas that are relatively flat in topography and focus imaging optics. Avoid excitation beam overspill when possible.
Operating Conditions | 9.2 | Recommendation | The following are suggested as a general guide for instrument operating conditions:
Operating Conditions | 9.2.1 | Recommendation | Turn on the X-ray source. Set the excitation voltage to at least 35 kV in order to provide sufficient overvoltage necessary for efficient X-ray excitation of the K-lines of higher atomic number elements, such as As, Rb, Sr, Zr, and Mo. Higher beam energies will typically improve detection limits of those elements. Once the beam excitation voltage is established, it should not be changed between specimens in a given comparison set. Most of the X-ray lines produced may be displayed with an energy range of 0 to 20 keV.
Operating Conditions | 9.2.2 | Recommendation | Select an appropriate X-ray optic size for the analysis of the specimen. For instruments with variable spot size options, this size is determined from the setpoint calibrated as described in 8.1.4. Once the spot size is established, it should not be changed between specimens in a given comparison set.
Operating Conditions | 9.2.3 | Recommendation | Set the pulse processor time constant at a midrange value; this is a compromise between maximum count rate and maximum spectral resolution. The optimal count rate is generally provided by the instrument manufacturer. Once established, do not change the pulse processor time between specimens in a given comparison set.
Operating Conditions | 9.2.4 | Recommendation | Adjust the beam current for each specimen as needed to yield a maximum X-ray detector dead time not to exceed 50 percent.
Specimen Analysis | 9.3 | Section Title
Specimen Analysis | 9.3.1 | Requirement | For each specimen, collect a spectrum for a live time that provides reasonable counting statistics for trace element peaks. For a µ-XRF system with a 100 µm monocapillary and a Si(Li) detector, 1200 live seconds is generally sufficient.
Specimen Analysis | 9.3.2 | Requirement | Collect replicate spectra to ensure that the questioned glass fragments and known glass source(s) are adequately characterized. When practical, analyze a minimum of three replicates on each questioned specimen examined and nine replicates on known glass sources.
Calculation And Interpretation Of Results | 10 | Section Title
Calculation And Interpretation Of Results | 10.1 | Requirement | Examine the spectrum, and identify and label the peaks.
Calculation And Interpretation Of Results | 10.2 | Requirement | Characteristic X-ray energies may be obtained from automatic element identification software; however, peak identifications shall be manually verified. Energy slide rules, computer generated theoretical fit curves that can be superimposed on the spectrum, and published tables can be used for this purpose.
Calculation And Interpretation Of Results | 10.3 | Recommendation | Spectral artifacts, such as sum peaks and escape peaks associated with the major peaks, should be considered and corrected.
Calculation And Interpretation Of Results | 10.4 | Requirement | In addition to the higher energy peaks, verify the presence of any lower energy peaks and their expected relative intensities. Individual asymmetric peaks and inconsistent peak ratios within a family may indicate a peak overlap.
Calculation And Interpretation Of Results | 10.5 | Informational Text | When the area of a characteristic energy of an element has a signal-to-noise ratio of three or more, that element may be identified. This criterion may not apply when spectral artifacts, overlapping energies of other elements, or shoulder peaks, or a combination thereof, are present within the characteristic energy range of that element.
Calculation And Interpretation Of Results | 10.6 | Requirement | Compare the spectra using peak identification, spectral comparisons, and peak intensity ratio comparisons.
Peak Identification | 10.6.1 | Requirement | Compare detected elements of the questioned and known glass spectra.
Spectral Comparisons | 10.6.2 | Requirement | Visually compare the spectral shapes and relative peak heights of the questioned and known glass specimen spectra.
Peak Intensity Ratio Comparisons | 10.6.3 | Requirement | When the area of a characteristic energy peak of an element has a signal-to-noise ratio of ten or more, that element may be used in semi-quantitative analysis such as peak intensity ratio comparisons. This criterion may not apply when spectral artifacts, overlapping energies of other elements, or shoulder peaks, or a combination thereof, are present within the characteristic energy range of that element. Evaluate elemental peak intensity ratios of the questioned and known glass spectra. Ratios for evaluation can include: Ca/Mg, Ca/Ti, Ca/Fe, Sr/Zr, Fe/Zr, and Ca/K, if those elements are present above the limit of quantitation (LOQ). These peak intensity ratio comparisons have been shown to provide the best discrimination among different sources of soda-lime glasses. Additional ratios should be chosen based on the elements present in the specimens. Improved precision may be obtained using ratios of peak intensities of elements close in atomic number because the effects of take-off angle are reduced. The elements present in the numerator and denominator of the ratio can be inverted if the denominator is a small number. Caution should be exercised when accepting background fits generated by automatic peak integration software because automatic background corrections are prone to errors. Therefore, manual adjustment of the background is recommended.
Interpretation of Comparisons | 10.7 | Section Title
Peak Identification | 10.7.1 | Recommendation | Reproducible differences in detected elements between specimens demonstrate that the specimens have different sources. When peak identification does not discriminate between the specimens, further spectral comparisons should be conducted.
Spectral Comparisons | 10.7.2 | Recommendation | Reproducible differences in spectral shapes and relative peak heights between specimens may indicate that the specimens have different sources. Peak intensity ratios can be calculated to demonstrate this difference. When evaluation of spectral shapes and relative peak heights do not discriminate between the specimens, peak intensity ratios should be calculated.
Peak Intensity Ratio Comparisons | 10.7.3 | Recommendation | Reproducible differences between specimens in peak intensity ratios can demonstrate that the specimens have different sources. One of the two following statistical measures is recommended to assess the association or discrimination of the samples based on elemental ratios:
Elemental Ratio Range Overlap | 10.7.3.1 | Recommendation | For each elemental ratio, compare the range of the questioned specimen replicates to the range for the known specimen replicates. Because standard deviations are not calculated, this statistical measure does not directly address the confidence level of an association. If the ranges of one or more elements in the questioned and known specimens do not overlap, it may be concluded that the specimens are not from the same source.
±3s | 10.7.3.2 | Recommendation | For each elemental ratio, compare the average ratio for the questioned specimen to the average ratio for the known specimens ±3s. This range corresponds to 99.7 % of a normally distributed population. If, for one or more elements, the average ratio in the questioned specimen does not fall within the average ratio for the known specimens ±3s, it may be concluded that the samples are not from the same source.
Precision And Bias | 11 | Section Title
Precision | 11.1 | Informational Text | An interlaboratory study was conducted in 2009 by the Elemental Analysis Working Group (EAWG) under National Institute of Justice grant 2009-DN-BX-K252. Standard reference glasses FGS 1 and FGS 2 of Latkoczy et al. (2005), which are soda-lime-silicate glasses manufactured by Schott (Schott, Germany) for the Bundeskriminalamt (BKA), were analyzed. Eight laboratories performed seven replicate analyses on full thickness fragments of the FGS glasses. A normalization of the data versus the CRM NIST SRM 1831 was conducted in order to standardize the responses from each laboratory and allow for direct comparison of the responses.
Precision And Bias | 11.2 | Informational Text | Results were reported based upon semi-quantitative analysis using elemental ratios. The following ratios were used for evaluation: Ca/Mg, Ca/Ti, Ca/Fe, Sr/Zr, Fe/Zr, and Ca/K. Element concentrations of NIST SRM 1831, FGS 1, and FGS 2 of Latkoczy et al. (2005), are presented in Appendix X1.
Precision And Bias | 11.3 | Informational Text | The precision data for FGS 1 and FGS 2 of Latkoczy et al. (2005) are tabulated in Appendix X2. The terms repeatability and reproducibility are used as specified in Practice E177. The 95 % limits were calculated by multiplying the respective standard deviations by a factor of 2.8.
Bias | 11.4 | Informational Text | No information can be presented on the bias of the procedure in this test method because this test method only allows for a semi-quantitative analysis of glass specimens. Relative areas of the characteristic X-ray energy peaks are compared between samples to arrive at a conclusion. The results of this test method are not directly comparable to known element concentrations.
Limit of Detection (LOD) and Limit of Quantitation (LOQ) | 12 | Section Title
Limit of Detection (LOD) | 12.1 | Informational Text | The LOD is equal to three times the concentration of an element in a standard divided by the signal-to-noise ratio of the peak of that element.
Limit of Quantitation (LOQ) | 12.2 | Informational Text | The LOQ is equal to ten times the concentration of an element in a standard divided by the signal-to-noise ratio of the peak of that element.
Limit of Detection (LOD) and Limit of Quantitation (LOQ) | 12.3 | Informational Text | LOD were determined for several elements in NIST SRM 1831, FGS 1, and FGS 2 of Latkoczy et al. (2005), from the data contained in the 2009 interlaboratory study (see 11.1). These results for µ-XRF systems used by nine participants in the EAWG study are presented in Table X3.1
Limit of Detection (LOD) and Limit of Quantitation (LOQ) | 12.4 | Informational Text | Signal-to-noise ratio calculations are described in Appendix X4.
Note on C6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: The cell is color-coded (fill) as listed below: o Section Title - Light Gray Fill o Requirement (e.g., “shall”) – No Fill o Recommendation (e.g., “should”) – Light Blue Fill o Informational Text (e.g., “can”) – Light Gray Fill If a clause, table or annex contains multiple categories of wording, the “highest rank” is chosen to determine the cell fill color. If the length of a clause exceeds the maximum row height, the clause type will include "cont'd".
Note on D6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Wording is entered as written in the standard and the cell is color-coded (fill) as listed below: o Section Title - Light Gray Fill o Requirement (e.g., “shall”) - No Fill o Recommendation (e.g., “should”) - Light Blue Fill o Informational Text (e.g., “can”) - Light Gray Fill If a clause, table or annex contains multiple categories of wording, the “highest rank” is chosen to determine the cell fill color. If the length of a clause exceeds the maximum row height, the clause wording will continue in the next row.
Note on E6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: List applicable management system documents and records that show implementation of/conformance to the Standard.
Note on F6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Columns G - J are set up as a Group that can collapse into a single column F. Column F will be automatically color coded to match the Implementation Status chosen in Column G to provide a quick visual overview of Implementation.
Note on G6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Implementation Status - Single choice; Standardized drop-down list; Cell fill color for visual & data filtering
Note on H6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Single choice; Standardized drop-down list that is dependent on the Implementation Status
Note on I6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field to document an implementation plan or other implementation notes (e.g., the reasons for no further implementation of only partial implementation or a conflicting law that will not allow implementation).
Note on J6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Format = MM/DD/YY
Note on K6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Columns L - Q are set up as a Group that can collapse into a single column K. Column K will be automatically color coded to match the Audit Status chosen in Column N to provide a quick visual overview of Auditing.
Note on L6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open Text field for the auditor to list the management system documents and records reviewed, examinations witnessed, interviews conducted during the audit that support the Audit Status chosen
Note on M6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field for the auditor to enter information related to their review of the documents and records listed in column K as well as information obtained during witnessing or interviews.
Note on N6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Single choice; Standardized drop-down list; Color-coded for visual & data sorting. oStatus Choices: Not Applicable: The requirement is not applicable to the scope of work being audited Document Review: A review of management system documents for conformance to requirements Conforming: Conforms with the requirement Opportunity for Improvement: Opportunity exists to improve a conforming practice Nonconforming: Does not conform to the requirement Resolved Nonconformity: Steps taken have resulted in conforming to the requirement
Note on O6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on the Audit Status choice; Only used when Audit Status = Opportunity for Improvement; Enter the wording of the Opportunity
Note on P6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on Audit Status choice; Only used when Audit Status = Nonconforming; Enter the wording of the Nonconformance
Note on Q6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Open text field; Cell fill color is dependent on Audit Status choice; Only used when original Audit Status = Nonconforming; Auditor enters information related to how the nonconformance was brought into conformance and changes the Audit Status to Resolved Nonconformity.

## Lists
Discipline | Clause Type | Implementation Status | Not Applicable | Will Not Implement | Not Yet Reviewed | Not Yet Implemented | Partial Implementation | Full Implementation | Audit Status
Anthropology | Section Title | Not Applicable | Not within scope of accreditation | Insufficient Resources | Evaluation of standard is pending | Insufficient Resources | Working to Full Implementation | Not Applicable | Not Applicable
Biology/DNA | Requirement | Will Not Implement | Method/Equipment not used | Conflicts with applicable law or regulation |  | Developing document | No Further Implementation Planned |  | Document Review
Bloodstain Pattern Analysis | Recommendation | Not Yet Reviewed |  | Disagrees with the provision |  | Requires method validation |  |  | Conforming
Crime Scene Investigation | Informational Text | Not Yet Implemented |  | Other |  | Other |  |  | Opportunity for Improvement
Digital Evidence | Requirement cont'd | Partial Implementation | [6 empty] | Nonconforming
Dogs & Sensors | Recommendation cont'd | Full Implementation | [6 empty] | Resolved Nonconformity
Facial Identification | Informational Text cont'd
Fire & Explosion Investigation
Fire Debris & Explosives
Firearms & Toolmarks
Footwear & Tire
Friction Ridge
Gunshot Residue Analysis
Interdisciplinary
Medicolegal Death Investigation
Odontology
Seized Drugs
Toxicology
Trace Evidence
Video/Imaging Technology & Analysis
Wildlife Forensics